DEFINITION Digital planimetry, is a measurement tool that provides a visible record of the dimensions of a wound. It enables the measurement of surface area, length, width, and amount of viable and non viable tissue. It can also be used to calculate the % surface area change, from the last wound area measurement.

SCOPE All clinical staff

CLINICAL ALERT Tracing Grid and depth measurement stick - Ensure that a fine point permanent marker pen is used to trace the wound. Not all pens are suitable to use on the grid. - Use minimal pressure when using a pen to trace the wound edge - Do not use the grid or depth measurement stick, if the seals are opened or damaged - Only insert the foam tipped end of the depth measurement stick, into the wound - Rest the measurement stick tip on the wound bed, using minimal pressure. Care should be taken not to depress the wound bed - Ensure that the top of the grid points toward the anatomical top of the wound (direction of the head of the patient) - Ensure that the patient details, wound position and date are documented in the relevant sections, on the grid

EQUIPMENT - Visitrak Digital Tablet - Visitrak Grid - Visitrak Depth Stick - Fine tipped permanent marker pen

PROCEDURE Symbols Guide Wound area in squared centimeters Length Width % change in wound area % non – viable tissue within the wound Wound depth in centimeter (using Visitrak Depth) To get started – Using the Visitrak tablet - Insert batteries. 2. Switch on. 3. Remove white backing paper from Visitrak Grid. 4. Place the grid area over the wound. Position the top of the grid toward the patients head. 5. Trace around the edge of the wound using a fine point permanent marker. 6. If the wound is too large to fit within the grid area of one tracing grid, overlap two grids and trace the wound across both grids. 7 Remove the contaminated wound contact layer from the tracing grid and dispose as clinical waste. 8. Place the tracing grid on the tablet – placing the holes over the pegs. If available, place the coversheet on top of the tracing grid and secure over the pegs on tablet. 9. Place the stylus at a point on the wound tracing and press the switch on the stylus and wait until the tracing symbol appears in the display. 10. Trace around the wound outline, without lifting the stylus. 11. The tablet will “beep” When the tracing is complete, and display the area measurement on the digital display. 11. If applicable, place the Visitrak depth into the deepest part of the wound and measure the depth using the scale on the stick. To measure the % change in the wound area After you have obtained the wound area measurement for the current wound – You can measure the % change in the wound surface area, from the previous wound. Press the symbol indicated and then use the key pad to enter the previous wound’s measurement. Press OK. The display will show the % change in the surface area (eg: this wound has decreased in size by 30%) To Calculate larger wounds using more than one grid After you have obtained the wound area measurement for the current wound- You can calculate the surface area of larger wounds, using more than one grid. Press +, as indicated and then measure the second wound by following steps 9 and 10 in this procedure. The tablet will beep and the total surface area of the wound will be displayed. To find the width: Press the symbol indicated To find the length: Press the symbol indicated To cancel a mistake or reset the process If you make a mistake, quickly press the C symbol. To reset it keep “C” pressed.

TROUBLE SHOOTING Replacing the Stylus Nib - Insert the nib of the stylus into the gap of the metal ring - Squeeze the metal ring to secure the nib and pull to remove - Place a new nib directly into the tip of the stylus and press firmly to secure Storage - Ensure the Visitrak Digital stylus is always returned to the tablet recess when not in use Cleaning Clean the Visitrak tablet, sylus and coversheet as per points listed below: - Wipe all device surfaces with 70% Isopropyl Alcohol Impregnated wipes eg: Rediwipe ® - Do not immerse any part of the Visitrak components in water - Do not sterilize the unit - Do not use plastic solvents or abrasive cleaners

EXPECTED OUTCOME - Wound healing will be optimized by effective and systematic monitoring of wound measurements. - Early identification of non responding wounds will ensure that treatment regimes can be re -evaluated and more appropriate interventions instigated, in a timely manner.

REFERENCES Visitrak, retrieved 10th November 2010 from http://wound.smith-nephew.com/au/Standard.asp?NodeId=4002

AUTHOR Lesley Stewart - Barwon South West Wound Management CNC

VALIDATION Victorian Regional Wound Management CNC’s